## Protocol
Single-Cell mRNA Seq PicoGreen Template
Number of Samples to Analyze: |  |  | Replicates
| 96 |  | 2
Lambda DNA Stock Concentration:
| 500 | ng/µL
Final well volume needed for Picogreen reader
25
1) Make/Get 1X TE buffer.
| Component | Vendor | Part Number | Lot Number | Vol. Required (µL)
| 20X TE Buffer | Invitrogen | Quant-iT Kit, P11496 |  | 501.024
| Ultra Pure Water | Invitrogen | 10977-023 |  | 9519.456
| TOTAL |  |  |  | 10020.48
2) Prepare diluted PicoGreen Working Solution.
| Component | Vendor | Part Number | Lot Number | Vol. Required (µL)
| 1X TE buffer |  |  |  | 4298.4
| PicoGreen Stock | Invitrogen | Quant-iT Kit, P11496 |  | 21.6
| TOTAL |  |  |  | 4320
3) Prepare 2 ng/µL Lambda DNA Solution.
| Component | Vendor | Part Number | Lot Number | Vol. Required (µL)
| Lambda Stock DNA | Invitrogen | Quant-iT Kit, P11496 |  | 2
| 1X TE buffer |  |  |  | 498
| TOTAL |  |  |  | 500
4) Prepare standard and background plate.
a) Label a 96-well plate as Standard Plate.
b) To the Standard Plate, pipette appropriate volumes of 1XTE as shown in Standard Plate Map below.
Note: Not added to A1.
d) Pipette appropriate volume 2 ng/µL Lambda DNA Solution into wells A1 and B1 of the Standard Plate, and make serial dilution by transferring from B1 to C1 as shown in Table 1 after pipetting up-down 10 times for thorough mixing. Be sure to remove volume from B2 so final volume is equal to all others. The table below is provided as an example. Alternatively, prepare 10 two-fold serial dilutions in strip tubes, vortexing well at each step. Concentrations are listed in table 1 below.
e) Pipette 2 µL C1 Harvest Reagent into well C1-C8 of the Standard Plate.
Table 1. Preparation of Standards and Background
Well | Conc. (pg/µL) λDNA | Volume of 1X TE (µL) | Volume of Diluent | Volume C1 Harvest
A1 | 2000 | 0 | 30 | -
B1 | 1000 | 30 | 30 | -
C1 | 500 | 30 | 30 | -
D1 | 250 | 30 | 30 | -
E1 | 125 | 30 | 30 | -
F1 | 62.5 | 30 | 30 | -
G1 | 31.25 | 30 | 30 | -
H1 | 15.625 | 30 | 30 | -
A1 | 7.8125 | 30 | 30 | -
B2 | 3.9062 | 30 | 30 | - | ***Remove and discard half of this dilution after mixing.
C2 | 1X TE Buffer | 30 | - | -
D2 | 1X TE Buffer | 30 | - | -
E2 | 1X TE Buffer | 30 | - | -
F2 | 1X TE Buffer | 30 | - | -
G2 | 1X TE Buffer | 30 | - | -
H2 | 1X TE Buffer | 30 | - | -
A3 | Harvest Reagent | 28 | - | 2
B3 | Harvest Reagent | 28 | - | 2
C3 | Harvest Reagent | 28 | - | 2
D3 | Harvest Reagent | 28 | - | 2
E3 | Harvest Reagent | 28 | - | 2
F3 | Harvest Reagent | 28 | - | 2
G3 | Harvest Reagent | 28 | - | 2
H3 | Harvest Reagent | 28 | - | 2
5) Sample dilution:
a) Label a 96-well plate as Sample Plate.
b) To each well of the Sample Plate, pipette 28 µl of 1XTE and 2 µl of Harvested C1 sample
d) Seal the plate with an adhesive film.
e) Vortex and spin down all components.
6) Pipette 30 ul diluted PicoGreen working solution into each well of the Standard Plate and Sample Plate.
a) Pipette down-up 5 times for mixing (or vortex and spin down).
7) Transfer 25 µl of the mix from both Standard Plate and Sample Plate to corresponding wells of a 384-well plate in duplicate as shown below.
384-well Fluorometer Plate
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
A | Standard 1 | Standard 1 | Standard 9 | Standard 9 | Harvest Reagent | Harvest Reagent
B | S1 | S1 | S9 | S9 | S17 | S17 | S25 | S25 | S33 | S33 | S41 | S41 | S49 | S49 | S57 | S57 | S65 | S65 | S73 | S73 | S81 | S81 | S89 | S89
C | Standard 2 | Standard 2 | Standard 10 | Standard 10 | Harvest Reagent | Harvest Reagent
D | S2 | S2 | S10 | S10 | S18 | S18 | S26 | S26 | S34 | S34 | S42 | S42 | S50 | S50 | S58 | S58 | S66 | S66 | S74 | S74 | S82 | S82 | S90 | S90
E | Standard 3 | Standard 3 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
F | S3 | S3 | S11 | S11 | S19 | S19 | S27 | S27 | S35 | S35 | S43 | S43 | S51 | S51 | S59 | S59 | S67 | S67 | S75 | S75 | S83 | S83 | S91 | S91
G | Standard 4 | Standard 4 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
H | S4 | S4 | S12 | S12 | S20 | S20 | S28 | S28 | S36 | S36 | S44 | S44 | S52 | S52 | S60 | S60 | S68 | S68 | S76 | S76 | S84 | S84 | S92 | S92
I | Standard 5 | Standard 5 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
J | S5 | S5 | S13 | S13 | S21 | S21 | S29 | S29 | S37 | S37 | S45 | S45 | S53 | S53 | S61 | S61 | S69 | S69 | S77 | S77 | S85 | S85 | S93 | S93
K | Standard 6 | Standard 6 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
L | S6 | S6 | S14 | S14 | S22 | S22 | S30 | S30 | S38 | S38 | S46 | S46 | S54 | S54 | S62 | S62 | S70 | S70 | S78 | S78 | S86 | S86 | S94 | S94
M | Standard 7 | Standard 7 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
N | S7 | S7 | S15 | S15 | S23 | S23 | S31 | S31 | S39 | S39 | S47 | S47 | S55 | S55 | S63 | S63 | S71 | S71 | S79 | S79 | S87 | S87 | S95 | S95
O | Standard 8 | Standard 8 | 1X TE | 1X TE | Harvest Reagent | Harvest Reagent
P | S8 | S8 | S16 | S16 | S24 | S24 | S32 | S32 | S40 | S40 | S48 | S48 | S56 | S56 | S64 | S64 | S72 | S72 | S80 | S80 | S88 | S88 | S96 | S96
10) Centrifuge the plate at 1,500 rpm for 1 minute.
11) Measure Fluorescence intensity on a 384-well Fluorometer.
12) Paste the raw data into the raw data sheet.
13) Check the results and sample dilution guide to determine the appropriate dilution factor for library preparation.
Fluidigm, the Fluidigm logo, and C1 are trademarks or registered trademarks of Fluidigm Corporation in the U.S. and/or other countries. All other trademarks are the sole property of their respective owners. For Research Use Only. Not for use in diagnostic procedures. © Fluidigm Corporation. All rights reserved.

## Raw Data
Paste your data starting in cell A-1.
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
A | 545697 | 576729 | 3244 | 3226 | 388 | 367 | 41 | 41 | 27 | 33 | 31 | 35 | 35 | 42 | 38 | 39 | 36 | 32 | 29 | 29 | 25 | 37 | 31 | 42
B | 69600 | 80071 | 4681 | 4912 | 57810 | 56129 | 73282 | 76260 | 84725 | 82891 | 88665 | 90131 | 49354 | 49347 | 3190 | 3214 | 89874 | 85738 | 2357 | 2276 | 33513 | 34901 | 96946 | 93448
C | 308292 | 357522 | 2074 | 1930 | 421 | 413 | 34 | 46 | 41 | 46 | 40 | 34 | 44 | 26 | 33 | 42 | 39 | 38 | 38 | 41 | 25 | 34 | 33 | 35
D | 65591 | 58812 | 77270 | 79953 | 81708 | 75184 | 44841 | 43989 | 81239 | 76881 | 3356 | 3421 | 6354 | 5981 | 45119 | 45159 | 3460 | 3263 | 90532 | 91079 | 65634 | 66151 | 82101 | 78750
E | 122551 | 168791 | 736 | 595 | 409 | 368 | 41 | 49 | 34 | 36 | 37 | 43 | 39 | 38 | 34 | 33 | 31 | 41 | 42 | 31 | 40 | 43 | 55 | 12
F | 79302 | 87913 | 3052 | 3479 | 55651 | 47013 | 83139 | 76976 | 79526 | 70671 | 61517 | 61533 | 64068 | 62544 | 3776 | 3613 | 83441 | 84853 | 54644 | 46459 | 3869 | 4357 | 77642 | 76119
G | 72819 | 90036 | 797 | 823 | 432 | 408 | 36 | 59 | 35 | 42 | 36 | 45 | 39 | 29 | 36 | 31 | 31 | 31 | 38 | 27 | 34 | 37 | 27 | 39
H | 95319 | 94260 | 88222 | 83287 | 65629 | 62655 | 65600 | 61763 | 76816 | 73388 | 54040 | 53024 | 3589 | 3484 | 64964 | 62225 | 85221 | 82732 | 70935 | 69685 | 3669 | 3689 | 60670 | 56260
I | 37187 | 42093 | 1320 | 811 | 433 | 447 | 34 | 32 | 50 | 40 | 32 | 38 | 36 | 40 | 25 | 36 | 19 | 38 | 34 | 32 | 30 | 32 | 39 | 31
J | 93954 | 94912 | 66639 | 63433 | 49792 | 44797 | 3871 | 3642 | 46674 | 44595 | 70862 | 69225 | 63706 | 66557 | 66771 | 60763 | 63496 | 63613 | 85652 | 84287 | 75448 | 76768 | 80643 | 73398
K | 18881 | 21869 | 655 | 647 | 414 | 416 | 34 | 42 | 29 | 26 | 37 | 45 | 37 | 34 | 41 | 38 | 42 | 38 | 27 | 39 | 30 | 42 | 48 | 42
L | 84084 | 85252 | 62391 | 61384 | 64585 | 61915 | 3874 | 3739 | 48955 | 46159 | 69234 | 64596 | 68288 | 67306 | 66580 | 63465 | 69998 | 69151 | 52126 | 49743 | 56765 | 59450 | 50108 | 48039
M | 10202 | 11298 | 761 | 693 | 425 | 439 | 41 | 36 | 35 | 35 | 37 | 34 | 35 | 38 | 44 | 43 | 43 | 43 | 33 | 35 | 28 | 40 | 32 | 36
N | 70755 | 65778 | 72934 | 72399 | 60259 | 59043 | 72026 | 69140 | 60677 | 59005 | 73534 | 71039 | 58203 | 59225 | 73677 | 70362 | 67822 | 68108 | 77088 | 74774 | 46798 | 46951 | 68709 | 66002
O | 5034 | 5635 | 805 | 791 | 483 | 465 | 40 | 54 | 29 | 36 | 40 | 37 | 40 | 40 | 30 | 49 | 31 | 45 | 41 | 43 | 37 | 30 | 49 | 44
P | 74914 | 60133 | 70901 | 69487 | 63728 | 62781 | 69669 | 66645 | 70287 | 73953 | 3806 | 3541 | 77980 | 81596 | 87470 | 84822 | 71848 | 70956 | 62598 | 61038 | 47041 | 47995 | 73715 | 71240

## Example Results
| Calibration Curve | [6 empty] | Harvest Reagent Background
| Std (ng/ul | A1 | A2 | Abg-Bkg |  |  |  |  | A1 | A2
| 2 | 545697 | 576729 | 560415 |  | Slope | 294276.3134 |  | 388 | 367
| 1 | 308292 | 357522 | 332109 |  | Intercept | 991.5782 |  | 421 | 413
| 0.5 | 122551 | 168791 | 144873 |  | Rsq | 0.9973 |  | 409 | 368
| 0.25 | 72819 | 90036 | 80629.5 |  |  |  |  | 432 | 408
| 0.125 | 37187 | 42093 | 38842 |  |  |  |  | 433 | 447
| 0.0625 | 18881 | 21869 | 19577 |  |  |  |  | 414 | 416
| 0.03125 | 10202 | 11298 | 9952 |  |  |  |  | 414 | 416
| 0.015625 | 5034 | 5635 | 4536.5 |  |  |  |  | 483 | 465
| 0.0078125 | 3244 | 3226 | 2437
| 0.0039062 | 2074 | 1930 | 1204 |  |  |  | Avgerage | 418.375
| 0 | 736 | 595 | 0 | 798
|  | 797 | 823
|  | 1320 | 811
|  | 655 | 647
Sample Read Intensity
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
A | 69600 | 80071 | 4681 | 4912 | 57810 | 56129 | 73282 | 76260 | 84725 | 82891 | 88665 | 90131 | 49354 | 49347 | 3190 | 3214 | 89874 | 85738 | 2357 | 2276 | 33513 | 34901 | 96946 | 93448
B | 65591 | 58812 | 77270 | 79953 | 81708 | 75184 | 44841 | 43989 | 81239 | 76881 | 3356 | 3421 | 6354 | 5981 | 45119 | 45159 | 3460 | 3263 | 90532 | 91079 | 65634 | 66151 | 82101 | 78750
C | 79302 | 87913 | 3052 | 3479 | 55651 | 47013 | 83139 | 76976 | 79526 | 70671 | 61517 | 61533 | 64068 | 62544 | 3776 | 3613 | 83441 | 84853 | 54644 | 46459 | 3869 | 4357 | 77642 | 76119
D | 95319 | 94260 | 88222 | 83287 | 65629 | 62655 | 65600 | 61763 | 76816 | 73388 | 54040 | 53024 | 3589 | 3484 | 64964 | 62225 | 85221 | 82732 | 70935 | 69685 | 3669 | 3689 | 60670 | 56260
E | 93954 | 94912 | 66639 | 63433 | 49792 | 44797 | 3871 | 3642 | 46674 | 44595 | 70862 | 69225 | 63706 | 66557 | 66771 | 60763 | 63496 | 63613 | 85652 | 84287 | 75448 | 76768 | 80643 | 73398
F | 84084 | 85252 | 62391 | 61384 | 64585 | 61915 | 3874 | 3739 | 48955 | 46159 | 69234 | 64596 | 68288 | 67306 | 66580 | 63465 | 69998 | 69151 | 52126 | 49743 | 56765 | 59450 | 50108 | 48039
G | 70755 | 65778 | 72934 | 72399 | 60259 | 59043 | 72026 | 69140 | 60677 | 59005 | 73534 | 71039 | 58203 | 59225 | 73677 | 70362 | 67822 | 68108 | 77088 | 74774 | 46798 | 46951 | 68709 | 66002
H | 74914 | 60133 | 70901 | 69487 | 63728 | 62781 | 69669 | 66645 | 70287 | 73953 | 3806 | 3541 | 77980 | 81596 | 87470 | 84822 | 71848 | 70956 | 62598 | 61038 | 47041 | 47995 | 73715 | 71240
Average Sample Read Intensity-Background
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 74417.125 | 4378.125 | 56551.125 | 74352.625 | 83389.625 | 88979.625 | 48932.125 | 2783.625 | 87387.625 | 1898.125 | 33788.625 | 94778.625
B | 61783.125 | 78193.125 | 78027.625 | 43996.625 | 78641.625 | 2970.125 | 5749.125 | 44720.625 | 2943.125 | 90387.125 | 65474.125 | 80007.125
C | 83189.125 | 2847.125 | 50913.625 | 79639.125 | 74680.125 | 61106.625 | 62887.625 | 3276.125 | 83728.625 | 50133.125 | 3694.625 | 76462.125
D | 94371.125 | 85336.125 | 63723.625 | 63263.125 | 74683.625 | 53113.625 | 3118.125 | 63176.125 | 83558.125 | 69891.625 | 3260.625 | 58046.625
E | 94014.625 | 64617.625 | 46876.125 | 3338.125 | 45216.125 | 69625.125 | 64713.125 | 63348.625 | 63136.125 | 84551.125 | 75689.625 | 76602.125
F | 84249.625 | 61469.125 | 62831.625 | 3388.125 | 47138.625 | 66496.625 | 67378.625 | 64604.125 | 69156.125 | 50516.125 | 59031.625 | 48655.125
G | 67848.125 | 72248.125 | 59232.625 | 70164.625 | 59422.625 | 71868.125 | 58295.625 | 71601.125 | 67546.625 | 75512.625 | 46532.625 | 66937.125
H | 67105.125 | 69775.625 | 62836.125 | 67738.625 | 71701.625 | 3255.125 | 79369.625 | 85727.625 | 70983.625 | 61399.625 | 47576.625 | 72059.125
Concentration Estimate (ng/ul)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 3.7427 | 0.17262 | 2.832 | 3.7394 | 4.2 | 4.485 | 2.4436 | 0.091345 | 4.4038 | 0.046209 | 1.6717 | 4.7806
B | 3.0987 | 3.9352 | 3.9267 | 2.1921 | 3.958 | 0.10085 | 0.2425 | 2.229 | 0.099475 | 4.5567 | 3.2868 | 4.0276
C | 4.1898 | 0.094582 | 2.5447 | 4.0089 | 3.7561 | 3.0642 | 3.155 | 0.11645 | 4.2173 | 2.5049 | 0.13778 | 3.8469
D | 4.7598 | 4.2993 | 3.1976 | 3.1741 | 3.7563 | 2.6568 | 0.1084 | 3.1697 | 4.2086 | 3.512 | 0.11566 | 2.9082
E | 4.7416 | 3.2432 | 2.3389 | 0.11961 | 2.2542 | 3.4984 | 3.248 | 3.1785 | 3.1677 | 4.2592 | 3.8075 | 3.8541
F | 4.2439 | 3.0827 | 3.1521 | 0.12216 | 2.3522 | 3.339 | 3.3839 | 3.2425 | 3.4745 | 2.5244 | 2.9584 | 2.4295
G | 3.4078 | 3.6321 | 2.9687 | 3.5259 | 2.9784 | 3.6128 | 2.9209 | 3.5991 | 3.3925 | 3.7985 | 2.3213 | 3.3614
H | 3.37 | 3.5061 | 3.1524 | 3.4023 | 3.6043 | 0.11538 | 3.9951 | 4.3192 | 3.5677 | 3.0791 | 2.3746 | 3.6225

## Sample Dilution Guide
Sample Dilution Guide
|  | 0.1-0.3 ng/ul (ideal) |  |  |  | <0.1 or >0.3 ng/ul (out of range) | [6 empty] | Number of samples at 0.1-0.3 ng/ul
1:15 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.24951 | 0.011508 | 0.1888 | 0.24929 | 0.28 | 0.299 | 0.16291 | 0.0060897 | 0.29359 | 0.0030806 | 0.11145 | 0.3187 | 78
B | 0.20658 | 0.26234 | 0.26178 | 0.14614 | 0.26387 | 0.0067234 | 0.016167 | 0.1486 | 0.0066317 | 0.30378 | 0.21912 | 0.26851
C | 0.27932 | 0.0063055 | 0.16964 | 0.26726 | 0.25041 | 0.20428 | 0.21033 | 0.0077633 | 0.28115 | 0.16699 | 0.0091854 | 0.25646
D | 0.31732 | 0.28662 | 0.21317 | 0.21161 | 0.25042 | 0.17712 | 0.0072264 | 0.21131 | 0.28057 | 0.23413 | 0.0077106 | 0.19388
E | 0.31611 | 0.21621 | 0.15592 | 0.007974 | 0.15028 | 0.23323 | 0.21654 | 0.2119 | 0.21118 | 0.28395 | 0.25384 | 0.25694
F | 0.28292 | 0.20551 | 0.21014 | 0.0081439 | 0.15682 | 0.2226 | 0.22559 | 0.21617 | 0.23163 | 0.16829 | 0.19723 | 0.16197
G | 0.22719 | 0.24214 | 0.19791 | 0.23506 | 0.19856 | 0.24085 | 0.19473 | 0.23994 | 0.22617 | 0.25323 | 0.15476 | 0.22409
H | 0.22466 | 0.23374 | 0.21016 | 0.22682 | 0.24028 | 0.0076919 | 0.26634 | 0.28795 | 0.23784 | 0.20528 | 0.1583 | 0.2415
1:14 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.26733 | 0.01233 | 0.20229 | 0.2671 | 0.3 | 0.32036 | 0.17455 | 0.0065247 | 0.31456 | 0.0033006 | 0.11941 | 0.34147 | 69
B | 0.22134 | 0.28108 | 0.28048 | 0.15658 | 0.28272 | 0.0072037 | 0.017322 | 0.15921 | 0.0071054 | 0.32548 | 0.23477 | 0.28769
C | 0.29927 | 0.0067558 | 0.18176 | 0.28635 | 0.26829 | 0.21887 | 0.22536 | 0.0083178 | 0.30124 | 0.17892 | 0.0098415 | 0.27478
D | 0.33998 | 0.30709 | 0.2284 | 0.22672 | 0.2683 | 0.18977 | 0.0077425 | 0.22641 | 0.30062 | 0.25086 | 0.0082614 | 0.20773
E | 0.33869 | 0.23166 | 0.16706 | 0.0085435 | 0.16102 | 0.24989 | 0.232 | 0.22704 | 0.22626 | 0.30423 | 0.27197 | 0.27529
F | 0.30313 | 0.22019 | 0.22515 | 0.0087256 | 0.16802 | 0.2385 | 0.24171 | 0.23161 | 0.24818 | 0.18031 | 0.21132 | 0.17354
G | 0.24342 | 0.25944 | 0.21205 | 0.25185 | 0.21274 | 0.25805 | 0.20864 | 0.25708 | 0.24232 | 0.27132 | 0.16581 | 0.2401
H | 0.24071 | 0.25044 | 0.22517 | 0.24302 | 0.25745 | 0.0082413 | 0.28537 | 0.30851 | 0.25483 | 0.21994 | 0.16961 | 0.25875
1:17 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.22016 | 0.010154 | 0.16659 | 0.21996 | 0.24706 | 0.26382 | 0.14374 | 0.0053732 | 0.25905 | 0.0027182 | 0.098338 | 0.28121 | 81
B | 0.18228 | 0.23148 | 0.23098 | 0.12895 | 0.23282 | 0.0059324 | 0.014265 | 0.13112 | 0.0058515 | 0.26804 | 0.19334 | 0.23692
C | 0.24646 | 0.0055636 | 0.14969 | 0.23582 | 0.22095 | 0.18025 | 0.18559 | 0.0068499 | 0.24808 | 0.14735 | 0.0081048 | 0.22629
D | 0.27999 | 0.2529 | 0.18809 | 0.18671 | 0.22096 | 0.15628 | 0.0063762 | 0.18645 | 0.24757 | 0.20659 | 0.0068035 | 0.17107
E | 0.27892 | 0.19078 | 0.13758 | 0.0070358 | 0.1326 | 0.20579 | 0.19106 | 0.18697 | 0.18633 | 0.25054 | 0.22397 | 0.22671
F | 0.24964 | 0.18133 | 0.18542 | 0.0071858 | 0.13837 | 0.19641 | 0.19905 | 0.19073 | 0.20438 | 0.14849 | 0.17403 | 0.14291
G | 0.20046 | 0.21365 | 0.17463 | 0.20741 | 0.1752 | 0.21251 | 0.17182 | 0.21171 | 0.19956 | 0.22344 | 0.13655 | 0.19773
H | 0.19823 | 0.20624 | 0.18543 | 0.20013 | 0.21202 | 0.006787 | 0.23501 | 0.25407 | 0.20986 | 0.18113 | 0.13968 | 0.21309
1:16 dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.23392 | 0.010789 | 0.177 | 0.23371 | 0.2625 | 0.28031 | 0.15273 | 0.0057091 | 0.27524 | 0.0028881 | 0.10448 | 0.29879 | 82
B | 0.19367 | 0.24595 | 0.24542 | 0.137 | 0.24738 | 0.0063032 | 0.015157 | 0.13931 | 0.0062172 | 0.28479 | 0.20543 | 0.25173
C | 0.26186 | 0.0059114 | 0.15904 | 0.25055 | 0.23476 | 0.19151 | 0.19719 | 0.0072781 | 0.26358 | 0.15655 | 0.0086113 | 0.24043
D | 0.29749 | 0.2687 | 0.19985 | 0.19838 | 0.23477 | 0.16605 | 0.0067747 | 0.19811 | 0.26304 | 0.2195 | 0.0072287 | 0.18176
E | 0.29635 | 0.2027 | 0.14618 | 0.0074756 | 0.14089 | 0.21865 | 0.203 | 0.19866 | 0.19798 | 0.2662 | 0.23797 | 0.24088
F | 0.26524 | 0.19267 | 0.19701 | 0.0076349 | 0.14701 | 0.20868 | 0.21149 | 0.20266 | 0.21716 | 0.15777 | 0.1849 | 0.15185
G | 0.21299 | 0.22701 | 0.18554 | 0.22037 | 0.18615 | 0.2258 | 0.18256 | 0.22495 | 0.21203 | 0.23741 | 0.14508 | 0.21009
H | 0.21062 | 0.21913 | 0.19702 | 0.21264 | 0.22527 | 0.0072112 | 0.2497 | 0.26995 | 0.22298 | 0.19245 | 0.14841 | 0.22641
1:6 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.62378 | 0.02877 | 0.472 | 0.62323 | 0.70001 | 0.7475 | 0.40727 | 0.015224 | 0.73397 | 0.0077015 | 0.27862 | 0.79676 | 1
B | 0.51645 | 0.65586 | 0.65445 | 0.36535 | 0.65967 | 0.016809 | 0.040417 | 0.3715 | 0.016579 | 0.75945 | 0.54781 | 0.67127
C | 0.6983 | 0.015764 | 0.42411 | 0.66814 | 0.62601 | 0.5107 | 0.52583 | 0.019408 | 0.70289 | 0.41748 | 0.022964 | 0.64115
D | 0.7933 | 0.71654 | 0.53293 | 0.52902 | 0.62604 | 0.4428 | 0.018066 | 0.52828 | 0.70144 | 0.58533 | 0.019276 | 0.48471
E | 0.79027 | 0.54053 | 0.38981 | 0.019935 | 0.37571 | 0.58307 | 0.54134 | 0.52975 | 0.52794 | 0.70987 | 0.63459 | 0.64234
F | 0.70731 | 0.51378 | 0.52536 | 0.02036 | 0.39204 | 0.55649 | 0.56399 | 0.54042 | 0.57909 | 0.42073 | 0.49307 | 0.40492
G | 0.56797 | 0.60535 | 0.49478 | 0.58765 | 0.4964 | 0.60213 | 0.48682 | 0.59986 | 0.56541 | 0.63309 | 0.38689 | 0.56023
H | 0.56166 | 0.58435 | 0.5254 | 0.56704 | 0.60071 | 0.01923 | 0.66585 | 0.71987 | 0.59461 | 0.51319 | 0.39576 | 0.60375
1:8 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.46784 | 0.021578 | 0.354 | 0.46742 | 0.525 | 0.56062 | 0.30546 | 0.011418 | 0.55048 | 0.0057761 | 0.20897 | 0.59757 | 8
B | 0.38734 | 0.49189 | 0.49084 | 0.27401 | 0.49475 | 0.012606 | 0.030313 | 0.27862 | 0.012434 | 0.56959 | 0.41085 | 0.50345
C | 0.52373 | 0.011823 | 0.31808 | 0.50111 | 0.46951 | 0.38303 | 0.39437 | 0.014556 | 0.52716 | 0.31311 | 0.017223 | 0.48087
D | 0.59497 | 0.53741 | 0.3997 | 0.39677 | 0.46953 | 0.3321 | 0.013549 | 0.39621 | 0.52608 | 0.439 | 0.014457 | 0.36353
E | 0.5927 | 0.4054 | 0.29236 | 0.014951 | 0.28178 | 0.4373 | 0.40601 | 0.39731 | 0.39596 | 0.5324 | 0.47594 | 0.48176
F | 0.53048 | 0.38534 | 0.39402 | 0.01527 | 0.29403 | 0.41737 | 0.42299 | 0.40531 | 0.43431 | 0.31555 | 0.36981 | 0.30369
G | 0.42598 | 0.45402 | 0.37109 | 0.44074 | 0.3723 | 0.45159 | 0.36512 | 0.44989 | 0.42406 | 0.47482 | 0.29017 | 0.42018
H | 0.42125 | 0.43826 | 0.39405 | 0.42528 | 0.45053 | 0.014422 | 0.49939 | 0.5399 | 0.44596 | 0.38489 | 0.29682 | 0.45281
1:10 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.37427 | 0.017262 | 0.2832 | 0.37394 | 0.42 | 0.4485 | 0.24436 | 0.0091345 | 0.44038 | 0.0046209 | 0.16717 | 0.47806 | 20
B | 0.30987 | 0.39352 | 0.39267 | 0.21921 | 0.3958 | 0.010085 | 0.02425 | 0.2229 | 0.0099475 | 0.45567 | 0.32868 | 0.40276
C | 0.41898 | 0.0094582 | 0.25447 | 0.40089 | 0.37561 | 0.30642 | 0.3155 | 0.011645 | 0.42173 | 0.25049 | 0.013778 | 0.38469
D | 0.47598 | 0.42993 | 0.31976 | 0.31741 | 0.37563 | 0.26568 | 0.01084 | 0.31697 | 0.42086 | 0.3512 | 0.011566 | 0.29082
E | 0.47416 | 0.32432 | 0.23389 | 0.011961 | 0.22542 | 0.34984 | 0.3248 | 0.31785 | 0.31677 | 0.42592 | 0.38075 | 0.38541
F | 0.42439 | 0.30827 | 0.31521 | 0.012216 | 0.23522 | 0.3339 | 0.33839 | 0.32425 | 0.34745 | 0.25244 | 0.29584 | 0.24295
G | 0.34078 | 0.36321 | 0.29687 | 0.35259 | 0.29784 | 0.36128 | 0.29209 | 0.35991 | 0.33925 | 0.37985 | 0.23213 | 0.33614
H | 0.337 | 0.35061 | 0.31524 | 0.34023 | 0.36043 | 0.011538 | 0.39951 | 0.43192 | 0.35677 | 0.30791 | 0.23746 | 0.36225
1:12 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.31189 | 0.014385 | 0.236 | 0.31162 | 0.35 | 0.37375 | 0.20364 | 0.0076121 | 0.36699 | 0.0038507 | 0.13931 | 0.39838 | 50
B | 0.25822 | 0.32793 | 0.32723 | 0.18267 | 0.32983 | 0.0084043 | 0.020209 | 0.18575 | 0.0082896 | 0.37973 | 0.2739 | 0.33564
C | 0.34915 | 0.0078818 | 0.21205 | 0.33407 | 0.31301 | 0.25535 | 0.26292 | 0.0097041 | 0.35144 | 0.20874 | 0.011482 | 0.32058
D | 0.39665 | 0.35827 | 0.26647 | 0.26451 | 0.31302 | 0.2214 | 0.009033 | 0.26414 | 0.35072 | 0.29267 | 0.0096382 | 0.24235
E | 0.39513 | 0.27026 | 0.1949 | 0.0099674 | 0.18785 | 0.29154 | 0.27067 | 0.26487 | 0.26397 | 0.35494 | 0.3173 | 0.32117
F | 0.35366 | 0.25689 | 0.26268 | 0.01018 | 0.19602 | 0.27825 | 0.28199 | 0.27021 | 0.28954 | 0.21037 | 0.24654 | 0.20246
G | 0.28399 | 0.30268 | 0.24739 | 0.29383 | 0.2482 | 0.30106 | 0.24341 | 0.29993 | 0.28271 | 0.31654 | 0.19345 | 0.28012
H | 0.28083 | 0.29217 | 0.2627 | 0.28352 | 0.30036 | 0.0096149 | 0.33293 | 0.35993 | 0.29731 | 0.2566 | 0.19788 | 0.30187
1:14 Dilution
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 0.26733 | 0.01233 | 0.20229 | 0.2671 | 0.3 | 0.32036 | 0.17455 | 0.0065247 | 0.31456 | 0.0033006 | 0.11941 | 0.34147 | 69
B | 0.22134 | 0.28108 | 0.28048 | 0.15658 | 0.28272 | 0.0072037 | 0.017322 | 0.15921 | 0.0071054 | 0.32548 | 0.23477 | 0.28769
C | 0.29927 | 0.0067558 | 0.18176 | 0.28635 | 0.26829 | 0.21887 | 0.22536 | 0.0083178 | 0.30124 | 0.17892 | 0.0098415 | 0.27478
D | 0.33998 | 0.30709 | 0.2284 | 0.22672 | 0.2683 | 0.18977 | 0.0077425 | 0.22641 | 0.30062 | 0.25086 | 0.0082614 | 0.20773
E | 0.33869 | 0.23166 | 0.16706 | 0.0085435 | 0.16102 | 0.24989 | 0.232 | 0.22704 | 0.22626 | 0.30423 | 0.27197 | 0.27529
F | 0.30313 | 0.22019 | 0.22515 | 0.0087256 | 0.16802 | 0.2385 | 0.24171 | 0.23161 | 0.24818 | 0.18031 | 0.21132 | 0.17354
G | 0.24342 | 0.25944 | 0.21205 | 0.25185 | 0.21274 | 0.25805 | 0.20864 | 0.25708 | 0.24232 | 0.27132 | 0.16581 | 0.2401
H | 0.24071 | 0.25044 | 0.22517 | 0.24302 | 0.25745 | 0.0082413 | 0.28537 | 0.30851 | 0.25483 | 0.21994 | 0.16961 | 0.25875